## Slide 1: ImageJ/FIJI image processing basics
Dr. Dominic Alibhai & Dr. Stephen Cross

## Slide 2: Overview
Brief Introduction to image acquisition and basic image processing techniques
Hands on image processing with Fiji/ImageJ
  Step by step guide to different image processing concepts in Fiji
  Exercises for you to complete on each section
Useful Websites for further reading:
  https://www.microscopyu.com/
  https://svi.nl/AboutMicroscopy
  http://zeiss-campus.magnet.fsu.edu/tutorials/
  http://www.leica-microsystems.com/science-lab/
  https://imagej.nih.gov/ij/docs/guide/user-guide.pdf

## Slide 3: Overview
Raw data + pre-processing
Image processing + segmentation
Object identification
Quantification

## Slide 4
Experimental Steps
Acquisition
Segmentation
Quantification
Sample prep
Visualisation
Notes: Good analysis requires good samples acquired properly
Crap in  crap out

## Slide 5
Resolution is the minimum distance between two points which allows them to be distinguished as individual objects
It is the NA of the objective which determines the resolution - not the magnification
D
F
θ
Objectives, Numerical Aperture and Resolution
Notes: In Camera Systems need to balance mag with NA as fixed camera pixel size to get best resolution

## Slide 6
Correction | Magnification | Numerical Aperture | F(trans) | F(epi)
Plan Achromat | 10x | 0.25 | 6.25 | 0.39
Plan Apo | 10x | 0.45 | 20.2 | 4.10
Plan Fluorite | 20x | 0.50 | 6.25 | 1.56
Plan Apo | 20x | 0.75 | 14.0 | 7.90
Plan Achromat | 40x | 0.65 | 2.64 | 1.11
Plan Apo | 40x (oil) | 1.30 | 11.0 | 18.0
Plan Fluorite | 60x | 0.85 | 2.01 | 1.45
Plan Apo | 60x (oil) | 1.40 | 5.4 | 10.6
Plan Apo | 100x (oil) | 1.40 | 1.96 | 3.84
Plan Apo | 100x (oil) | 1.45 | 2.10 | 4.42
For a given NA, lower magnification gives a brighter image
Magnification and image brightness
Notes: Camera systems -> need to balance light throughput with pixel size

## Slide 7: Bit Depth and digital images
20 | 24 | 31 | 38 | 39 | 43 | 40 | 39 | 41 | 39 | 45 | 44 | 43 | 40 | 35 | 26 | 23 | 21 | 16 | 15 | 12 | 15 | 10 | 9
24 | 27 | 31 | 33 | 39 | 36 | 41 | 38 | 39 | 40 | 42 | 45 | 43 | 44 | 34 | 29 | 27 | 22 | 17 | 14 | 13 | 11 | 8 | 10
25 | 25 | 31 | 37 | 41 | 41 | 41 | 41 | 37 | 36 | 36 | 45 | 43 | 40 | 39 | 33 | 27 | 23 | 15 | 16 | 15 | 13 | 12 | 13
20 | 25 | 24 | 36 | 39 | 38 | 41 | 37 | 39 | 39 | 38 | 39 | 43 | 37 | 39 | 35 | 30 | 24 | 20 | 19 | 12 | 14 | 13 | 11
24 | 25 | 29 | 36 | 35 | 39 | 37 | 39 | 38 | 37 | 38 | 42 | 41 | 42 | 42 | 38 | 37 | 27 | 26 | 23 | 19 | 17 | 17 | 15
27 | 24 | 32 | 38 | 37 | 42 | 37 | 39 | 41 | 41 | 35 | 41 | 38 | 48 | 42 | 40 | 36 | 31 | 29 | 26 | 20 | 20 | 17 | 19
26 | 25 | 31 | 32 | 32 | 38 | 40 | 37 | 38 | 35 | 40 | 39 | 40 | 46 | 40 | 39 | 38 | 41 | 35 | 28 | 27 | 24 | 19 | 20
27 | 22 | 32 | 35 | 38 | 34 | 38 | 40 | 37 | 39 | 40 | 41 | 37 | 37 | 38 | 43 | 43 | 41 | 38 | 36 | 31 | 30 | 27 | 25
24 | 30 | 30 | 40 | 40 | 40 | 39 | 36 | 41 | 34 | 39 | 40 | 40 | 40 | 41 | 42 | 42 | 44 | 42 | 42 | 37 | 36 | 33 | 37
28 | 34 | 37 | 41 | 41 | 40 | 43 | 38 | 41 | 42 | 40 | 37 | 35 | 35 | 40 | 37 | 39 | 41 | 42 | 41 | 42 | 42 | 40 | 40
31 | 35 | 41 | 40 | 42 | 44 | 45 | 46 | 43 | 37 | 42 | 39 | 36 | 39 | 39 | 38 | 41 | 39 | 40 | 38 | 45 | 42 | 42 | 41
34 | 38 | 40 | 42 | 41 | 44 | 44 | 48 | 47 | 41 | 43 | 38 | 39 | 39 | 37 | 39 | 39 | 39 | 39 | 37 | 42 | 44 | 40 | 39
38 | 41 | 41 | 47 | 46 | 49 | 46 | 44 | 47 | 42 | 39 | 41 | 36 | 39 | 36 | 38 | 36 | 35 | 39 | 37 | 37 | 39 | 42 | 41
42 | 48 | 47 | 45 | 50 | 49 | 46 | 46 | 46 | 50 | 48 | 47 | 40 | 35 | 36 | 40 | 37 | 39 | 36 | 40 | 37 | 36 | 37 | 41
44 | 45 | 54 | 52 | 54 | 48 | 50 | 45 | 47 | 44 | 47 | 43 | 50 | 41 | 41 | 36 | 38 | 36 | 41 | 39 | 37 | 37 | 37 | 37
44 | 47 | 47 | 55 | 49 | 50 | 48 | 51 | 46 | 44 | 44 | 45 | 47 | 45 | 40 | 43 | 36 | 40 | 36 | 36 | 36 | 32 | 37 | 39
43 | 46 | 50 | 51 | 54 | 52 | 50 | 48 | 47 | 47 | 51 | 46 | 42 | 41 | 42 | 41 | 41 | 39 | 38 | 38 | 38 | 39 | 38 | 40
45 | 50 | 50 | 47 | 43 | 48 | 44 | 45 | 50 | 45 | 45 | 44 | 48 | 44 | 43 | 43 | 43 | 40 | 41 | 34 | 35 | 38 | 40 | 38
46 | 45 | 43 | 42 | 34 | 38 | 36 | 41 | 42 | 43 | 44 | 43 | 42 | 44 | 51 | 48 | 43 | 41 | 41 | 38 | 36 | 39 | 35 | 39
41 | 39 | 39 | 35 | 35 | 29 | 29 | 32 | 33 | 36 | 39 | 37 | 44 | 43 | 46 | 43 | 43 | 43 | 39 | 40 | 39 | 39 | 41 | 39
38 | 30 | 34 | 29 | 28 | 29 | 27 | 26 | 28 | 32 | 33 | 38 | 42 | 42 | 44 | 44 | 45 | 48 | 39 | 43 | 40 | 45 | 39 | 43
29 | 30 | 29 | 24 | 24 | 23 | 27 | 29 | 25 | 28 | 30 | 38 | 33 | 40 | 43 | 43 | 46 | 44 | 46 | 46 | 42 | 42 | 46 | 42
26 | 24 | 25 | 23 | 22 | 21 | 26 | 25 | 23 | 28 | 28 | 29 | 35 | 41 | 39 | 42 | 41 | 43 | 48 | 48 | 44 | 45 | 46 | 40
25 | 24 | 22 | 22 | 20 | 22 | 17 | 19 | 23 | 23 | 25 | 27 | 32 | 36 | 40 | 41 | 49 | 46 | 40 | 48 | 47 | 43 | 48 | 44
1 bit | 21 | 2
8 bit | 28 | 256
12 bit | 212 | 4096
16 bit | 216 | 65536

## Slide 8: RGB Images
RGB images are made up of 3 x 8-bit images - One ‘Red’, one ‘Green’ and one ‘Blue’
An RGB image therefore contains 3 values per pixel representing the intensity of each of the 3 colours.
For analysis, RGB images can be split into the constituent channels, or kept as RGB depending on the analysis performed.
Red
Green
Blue

## Slide 9
Oversampled
Not ‘bad’ but will take longer and cause more photodamage to the sample
Sampling an image
Nyquist sampled
Minimum sampling rate to capture the signal
Under sampled
Aliasing
Notes: Sampling of image just like sampling other things e.g. converting music from analogue to digital

## Slide 10: Sampling an image
Smaller features in image correspond to higher spatial frequencies
When setting pixel size (in confocal), important to sample image appropriately to avoid aliasing in image.
Rule of Thumb:
Not always possible due to photobleaching etc.
Need to be aware of roughly the smallest feature you are interested in.
http://svi.nl/NyquistRate - lots of information on Nyquist sampling + Nyquist calculator! http://svi.nl/NyquistCalculator
Notes: Recommend registering with SVI for more useful info
SVI uses frequency bandwidth to calculate nyquist rate (fourier transform of psf =optical transfer function) which is smaller than ‘rule of thumb’ -> for airy1 pinhole can use sampling slightly larger than svi recommendation

## Slide 11: Image Processing
Improve the signal to noise of the image
  Background subtraction
  Filtering
Segment and measure features within the image
  Size
  Shape - Circularity/ellipticity etc.
  Intensity – per channel or ratio
  Tracking
  Changes over time
  Changes after stimulation etc.

## Slide 12: File Types
TIFF – good general choice
  Lossless storage
  Header tags for metadata
  Multipage compatible (i.e. stacks, multicolour, tile scanning etc.)
Proprietary formats (e.g. .lif) – most complete
  Lossless storage
  (Should!) Contain all metadata
  Readers required for viewing outside proprietary software
PNG – good for transfer/display
  Lossless compression
  RGB only
  No associated metadata
JPEG – ….just don’t…

## Slide 13: Image Types
2D
Hyperstacking (4,5,6,7…D)
XY
XYT
XYC
XYZ
3D
XYZT
XYZCT

## Slide 14: Histograms & Thresholds
Histogram of image shows probability density function of intensity values
  Range of intensity values determined by bit depth of image
Thresholding separates foreground (objects of interest) from background
Various algorithms available in FIJI:
  Otsu (default)
  Li
  Percentile
  …
Each will give different results
  Doesn’t really matter which one you pick!
255
0
Frequency
Intensity
Threshold

## Slide 15: Binary Images
1 bit | 21 | 2
8 bit | 28 | 256
12 bit | 212 | 4096
16 bit | 216 | 65536
Binary image formed of 0’s and 1’s – background and objects
Use threshold value calculated from histogram to determine cut-off intensity
  Otsu threshold used – threshold = 23

## Slide 16: Filtering Images
Spatial Filter:
  Image operation where each pixel value I(u,v) is changed by a function of the intensities of pixels in the neighbourhood of (u,v).
  Example: Mean filter in 3 x 3 neighbourhood
I
u
v
I(u,v)
I’
u
v
I’(u,v)
u
u+1
u-1
v
v+1
v-1

## Slide 17: Filtering Images
Applying filters:
  Raster scan processing kernel across image
1
1
1
1
1
1
1
1
1
3 x 3 Mean kernel

## Slide 18: Filtering images
No Filter
Gaussian (r=2)
Median (r=2)
Sobel Edge Detection
Notes: Although I use jpeg in this course!

## Slide 20: ImageJ/Fiji interface
File open/save
Copy/paste, invert image intensity, regions of interest
Stack operations, scale, crop, zoom, rename, LUTs
Filter, binarise, add/subtract images, subtract background
Measurement, analyse particles, intensity profiles, colocalisation
Region selection
Zoom/pan
LUTs
Downloaded plugins (many installed with Fiji)
System messages, pixel coordinates, pixel intensity

## Slide 21: Loading images
Load standard images using “Open”
  File > Open
To load image sequences
  Files ending in “0001.tif”, “0002.tif”, etc.
  File > Import > Image Sequence…
Load Leica (.lif) files using Bio-Formats
  File > Import > Bio-Formats
  Included with FIJI (download for ImageJ)

## Slide 22: Loading images
Display OME-XML metadata
  Shows formatted metadata
  Microscope and image acquisition settings (laser power, objective lens, etc.)
Use virtual stack
  Images loaded directly from storage (not loaded to RAM in advance)
  Useful for large files (>1GB)

## Slide 23: Loading images
Select relevant images
  Only shows up when >1 image stack in file

## Slide 24: Image metadata
OME-XML metadata
  Loaded via Bio-Formats
  Easier to read
Show info
  Image > Show Info…
  Plain text format
  Works for any image (content will vary)
  Shows more information than OME-XML

## Slide 25: Saving images
Save images to single .tif file
  File > Save As > Tiff…
Save image sequence
  File > Save As > Image Sequence…
  Each slice saved as a separate .tif file
  Filenames end with “0001.tif”, “0002.tif”, etc.
  Number of zeroes dictated by “Digits” option

## Slide 26: Image navigation
Move through frames using sliders
  Channel (C)
    Blue, red, etc.
  Z-slice (Z)
  Time ()
    Click  to play video
Change video speed
  Image > Stacks > Animation > Animation options
Thanks to Joe Wrennall for confocal images

## Slide 27: Worksheet – section 1
Aims
  Loading an image with BioFormats
  Selecting channels and slices with sliders
Notes
  Download “ExampleData.lif” from https://goo.gl/YrBvRU
Notes: 45 minutes
Wait 5 minutes
15:00

## Slide 28: Worksheet – section 1

## Slide 29: Image navigation
Zooming
  Select
  Left click to zoom in, right click to zoom out
  Can also zoom with up/down keys on keyboard
Dragging
  Select
  Click and hold to drag
  Position in image shown in top-left
Zoom
Drag

## Slide 30: Image navigation
Pixel coordinates
  Move the cursor over the image
  Coordinates and intensity displayed in main ImageJ/FIJI window
  If distance per pixel known, position shown in that scale, otherwise shown in pixel units
  Calibration automatically loaded if available
Manually setting distance per pixel
  Analyze > Set Scale…

## Slide 31: Brightness and contrast
Adjust how intensity is displayed
  Image > Adjust > Brightness/Contrast…
  By default 0 intensity is displayed as black and 255 (8-bit) as white
  Use min and max sliders to change display range
  Image intensity values are unchanged
Alter intensity values (optional)
  Clicking “Apply” rescales intensity values to 0-255
  Treat this function with care
Notes: It’s this reason 16-bit images can look entirely black when opened in Windows.

## Slide 32: Lookup table (LUT)
Microscope images generally greyscale
  No colour information
  Colour determined by LUT
  Image data not altered by LUT

## Slide 33: Worksheet – section 2
Aims
  Adjusting appearance of image with Brightness/Contrast
  Changing lookup table (LUTs)
  Using zoom and pan tools
Notes: 45 minutes
Wait 5 minutes
15:00

## Slide 34: Worksheet – section 2

## Slide 35: Stack manipulation
Split channels
  Image > Color > Split channels
Merge channels
  Image > Color > Merge channels…
Duplicate stack
  Image > Duplicate
  Can duplicate individual slices or whole stack

## Slide 36: Stack manipulation
Rename stack
  Image > Rename
Extract single slices from stack
  Image > Stacks > Tools > Make Substack…
  Enter single channel/slice numbers or ranges

## Slide 37: Intensity projections
Project stack to single plane
  Image > Stacks > Z Project…
  Max Intensity: Highlights signal
  Min intensity: Good for identifying constant background
  Also available: mean, median, sum and standard deviation
Max intensity
Min intensity
Mean intensity
Intensity standard dev.

## Slide 38: Worksheet – section 3
Aims
  Split two channels of input image
    Discard the cell image (channel 2)
  Compare approaches for extracting “good example images of nuclei”
    One image from a Z-projection
    One image from a single Z-slice
Notes: 45 minutes
Wait 5 minutes
15:00

## Slide 39: Worksheet – section 3
Substack
Max. intensity Z-projection

## Slide 40: Background subtraction
Rolling ball
  Passes a ball underneath the image
  The ball traces out features larger than its own diameter (i.e. the background)
  Repeat from opposite side to remove effect of ball radius

## Slide 41: Background subtraction
Rolling ball
  Process > Subtract Background…
  Tick “Create background” to get background image
  “Sliding paraboloid” replaces the ball with a parabola and can give better results
Rolling ball
Sliding paraboloid

## Slide 42: Image calculator
Add/subtract images
  Process > Image Calculator…
  Calculated pixel-by-pixel
  “32-bit (float) result” when negative or non-integer values may occur
Image B (background)
Image C (result)
Image A (signal)

## Slide 43: Worksheet – section 4
Aims
  Use two approaches to subtract background
    Rolling ball subtraction
    Minimum intensity projection subtraction
  Apply both subtraction methods to nuclei image from exercise 3
    Duplicate image first, then apply subtraction
    End up with two images for comparison

## Slide 44: Worksheet – section 4
Rolling ball subtraction
Minimum intensity Z-projection subtraction

## Slide 45: Selecting regions
Hand-drawing tools to select regions
  Right click for more options (red arrow)
  Can drag handles to reshape afterwards
Polygon selection
  Close loop to finish (small square)
Line tools
  Don’t need to enclose area

## Slide 46: Selecting regions
Automatic region selection using wand tool
  Selects contiguous pixel regions with similar intensities
  Double click icon to get more options
    Tolerance sets permitted intensity variation

## Slide 47: Making measurements
Setting measurements
  Analyze > Set Measurements…
  “Redirect to” uses the regions from a binary image on a different image
Making measurements
  Analyze > Measure

## Slide 48: Results table
Lists measurement output
  One measurement per row
Extra options when window is called “Results”
  “Summarize” gives mean, SD, etc. of all columns
  “Distribution” creates histogram of one column
  “Sort” organises data based on one column
  “Plot” creates XY plot of two columns

## Slide 49: Linear intensity profile
Draw line on image
  Straight, freehand, segmented, etc.
Measure line profile
  Analyze > Plot Profile
  Profile displayed in new window

## Slide 50: Region Of Interest (ROI) manager
Store selected regions in ROI manager
  Add ROIs with Edit > Selection > Add to Manager
  Simultaneously measure all ROIs (“Measure” button)
Transfer ROIs to another image
  Select other image, then click ROI in list
Save and load ROIs to file from “More >>”
  .roi for single ROI, .zip for multiple ROIs

## Slide 51: Worksheet – section 5
Aims
  Enable specific measurements (“Area” in this case)
  Select and add regions to the ROI manager
  Measure all ROIs simultaneously
  Calculate the mean ROI (nucleus) area using the Results Table “Summarize” function

## Slide 52: Worksheet – section 5

## Slide 53: Image filtering
Apply spatial filter
  Process > Filters > [multiple]
  Sobel edge filter at Process > Find Edges
  Many require range for filtering
Custom spatial filter
  Process > Filters > Convolve…
  Specify filter kernel
Unfiltered
1px Gaussian
2px median
2px Gaussian
Edge (Sobel)
5px Gaussian

## Slide 54: Image filtering
Smooth (mean filter σ = 1px)
  Process > Smooth
Despeckle (median filter σ = 1px)
  Process > Noise > Despeckle
Remove outliers
  Process > Noise > Remove Outliers…
  Pixels deviating from local median by more than “Threshold” replaced by local median
“Raw”
Radius = 1px
Radius = 2px

## Slide 55: Worksheet – section 6
Aims
  Test different image filters to reduce noise
  We wish to remove speckle noise whilst retaining sharp edges for nuclei

## Slide 56: Worksheet – section 6
Input image
1px median filter

## Slide 57: Intensity thresholding
Manual intensity thresholds
  Image > Adjust > Threshold…
  Set percentage above threshold
5% over threshold
15% over threshold
25% over threshold

## Slide 58: Intensity thresholding
Automatic global intensity thresholds
  Image > Adjust > Auto Threshold
  Uses a single value for the entire image
  Ability to test multiple algorithms
    http://imagej.net/Auto_Threshold
Automatic local intensity thresholds
  Image > Adjust > Auto Local Threshold
  Set size of local region

## Slide 59: Intensity thresholding

## Slide 60: Worksheet – section 7
Aims
  Test different image threshold methods to binarise image
  The main aim is to have well-defined nuclei
    A few holes in nuclei or speckles in the background aren’t a problem
  End up with black nuclei on a white background

## Slide 61: Worksheet – section 7
Input image
Huang threshold
Invert

## Slide 62: Binary operations
Make binary image
  Process > Binary > Make Binary
  Apply to intensity thresholded image
Binary operations
  Process > Binary > [multiple]
Binary options
  Process > Binary > Options
  Iterations, operation preview, etc.
Dilate
Outline
Erode
Skeletonize
Fill holes
Distance map
Unaltered

## Slide 63: Watershed segmentation
Apply binary watershed
  Process > Binary > Watershed
  Useful for splitting touching objects
  Use prior to particle analysis
  Calculates distance map, then grows points from minima until touching (object boundary)
Before watershed
After watershed

## Slide 64: Worksheet – section 8
Aims
  Apply binary filters to clean up image
    Remove holes in nuclei
    Split nuclei in contact
  Speckles in background still aren’t a problem
    We’ll deal with these later

## Slide 65: Worksheet – section 8
Input image
Fill holes (?)
Watershed transform

## Slide 66: Analyse particles
Run analyse particles
  Analyze > Analyze Particles…
Set limits on particles
  Size
    Image scale units if set, option for pixels
  Circularity
    Non-circular (0) or perfectly circular (1)
  Exclude on edges:
    Remove particles touching image border
Overlay (default)
Counts (“grey” LUT)
Counts (“glasbey” LUT)

## Slide 67: Worksheet – section 9
Aims
  Detect nuclei using Analyze Particles tool
  Redirect measurements to “best” nuclei image
    Analyze > Set Measurements
  Display identified objects as outlines or “count masks”

## Slide 68: Worksheet – section 9
Input image
Count masks

## Slide 69: Plotting data
Plot data from Results Table
  Only works if Results Table is called “Results”
XY plot of two columns
  (From Results Table) Results > Plot…
  “More” button to set plotting options
    Label sizes, legend, export high resolution
Histogram of one column
  (From Results Table) Results > Distribution…

## Slide 70: Fitting data
Fit curves to XY plots
  (From plot window) Data > Add Fit…
Multiple curve options
  Including straight line, polynomial, Gaussian and exponential
Fitting parameters displayed in Log

## Slide 71: Worksheet – section 10
Aims
  Plot histogram of nuclear area measurements
    5 bins from 100 to 200
  Plot “Bin start” vs “Count” as XY plot
    Will need to rename new results table as “Results”
  Fit Gaussian and get centre and standard deviation

## Slide 72: Worksheet – section 10

## Slide 73: Installing plugins
Extend functionality with extra plugins
Go to Help > Update…
  Select “Manage update sites”
  Tick the plugin packages to be installed
  Or, add a new update site
  Click “Apply changes” to complete
Notes: When running

## Slide 74: Macro recording
Allows short, simple bits of code to be automated
  Plugins > Macros > Record…
  Will turn each mouse click into macro code
  Create button sends code to macro editor

## Slide 75: “Batching” Macros
Process > Batch > Macro…
  Set input containing folder of images to be processed
  Set output for output images
  Either:
    Copy code from recorder
    Open saved macro

## Slide 76
https://github.com/wbif-bristol/
ImageJ-Fiji-Level-2-course
https://github.com/wbif-bristol/
CellProfiler-course